Assessing variability and uncertainty in Radiofrequency Electromagnetic Field Exposure measurements using the ETAIN 5G Scientist App

Author

Berend den Boer

Scientific Report

Faculty of Veterinary Medicine, University of Utrecht

IRAS (Institute for Risk Assessment Sciences)

ETAIN (Exposure to Electromagnetic Fields and Planetary Health)

Examiner: Dr. A. (Anke) Huss

Daily Supervisor: Lea Belácková

05-07-2025

Plain Language Summary

As more people use mobile phones and 5G networks grow, there is also more concern about exposure to radiation from antennas. This radiation is called radiofrequency electromagnetic fields (RF-EMF). To help people understand how much of this radiation they are exposed to, researchers developed the ETAIN 5G Scientist App. This smartphone app lets anyone measure RF-EMF levels around them and could help build a map of exposure levels across Europe.

In this research project, we tested how accurate and reliable the app is. We compared the app’s measurements to those of a professional, scientific device called the ExpoM-RF 4, which is considered a trusted tool for measuring RF-EMF. We did this in real-world situations both in a static position and while moving around (mobile). This has been done across different places (urban vs. rural, indoor vs. outdoor) and times of day.

We found that the app generally noticed the same patterns as the professional device: for example, RF-EMF levels were higher outdoors than indoors, and higher in cities than in little villages. Exposure was also higher during the day than at night, which makes sense because more people use mobile data during the day. However, the app was more variable in its readings, especially when signal levels were high.

In the static measurements, the app usually overestimated exposure, often reporting values up to 10 times higher than the professional device. In moving (mobile) measurements, the app sometimes underestimated exposure, especially in crowded areas. The app compared to the professional device showed similar trends overall.

This study shows that the app is a useful tool for raising awareness about RF-EMF exposure and could help with citizen science projects. It is not as accurate as scientific equipment yet, especially environments where exposure levels are high. Still, the app has great potential. It can give people insight into their everyday exposure and help scientists collect large amounts of data across Europe.

Abstract

The public interest in radiofrequency electromagnetic fields (RF-EMF), especially due to the change from 4G to 5G networks, has generated demand for accessible tools to assess RF-EMF exposure. The ETAIN 5G Scientist App aims to enable citizen science by allowing individuals to estimate downlink RF-EMF exposure using smartphones. With these measurements, the ETAIN research group want to eventually map the exposure levels in Europe. This study evaluates the variability and uncertainty of the app by comparing them to those from a calibrated reference device, the ExpoM-RF 4, under real-world conditions.

Measurements were conducted using both static and mobile approaches across a range of environmental and temporal conditions. Static (24-hour) and mobile (15-minute) RF-EMF measurements were carried out throughout different locations in the Netherlands using three smartphones equipped with SIM cards from Vodafone, KPN, and Odido, alongside the ExpoM-RF 4. Exposure data were converted to power flux density (mW/m²), and nonparametric statistics, including Wilcoxon signed-rank tests, were used to assess differences in exposure levels.

Results show that the ETAIN 5G Scientist App generally followed the same exposure trends as the ExpoM-RF 4, detecting higher exposure in outdoor versus indoor locations and in urban versus rural settings. Temporal patterns were also observed, with higher exposure during the day than at night. However, the app displayed greater variance, particularly under higher demand of the network and exposure conditions.

Across all static conditions, the app significantly overestimated RF-EMF exposure levels compared to the ExpoM-RF 4, often by a factor of ten. In mobile settings, the app occasionally underestimated exposure, particularly in high-demand environments. The strength of correlation between the app and the reference device was moderate overall (R = 0.76 p < 2.2e−16 for static, R = 0.51 p < 2.2e−16 for mobile).

This research highlights both the potential and the current limitations of smartphone-based RF-EMF monitoring. While the ETAIN 5G Scientist App can reflect overall trends, its calibration requires refinement especially for high exposure environments, before it can reliably estimate exposure levels. Nevertheless, its accessibility makes it a promising tool for public engagement and large-scale exposure mapping through citizen science.

- Table of Contents

1	Introduction	5

2	Methods	8

2.1	Study Design	8

2.2	Materials	10

2.3 Measurement Setup and Procedure	12

2.4 Data handling	13

2.5 Variability analysis	14

2.6 Uncertainty analysis	14

3	Results	16

3.1 Variability Static measurements	16

3.2 Variability Mobile measurements	21

3.3 Uncertainty Static measurements	25

3.4 Uncertainty Mobile measurements	30

4	Conclusion and Discussion	33

5	GenAI disclosure statement	36

6	References	37

7	Appendix	38

Appendix A	Measurement details and procedure	38

Appendix B	Conversion Data ETAIN 5G Scientist app	41

Appendix C	Variability analysis for each location	43

Appendix D	Uncertainty analysis for each location	79

Appendix E	Tables comparison Static and Mobile measurements	102

# Introduction

Changes in wireless telecommunication technologies and electronic devices have raised concerns about possible changes in exposure to radiofrequency electromagnetic fields (RF-EMFs). RF-EMFs are a type of non-ionizing radiation emitted by mobile phones, radio’s, WIFI routers and especially antennas from mobile network base stations. In recent years, Long-Term Evolution (LTE/4G) cellular network technology is being used for wireless communication and at the same time, new generation (5G/NR) is being implemented.

Compared to previous generations such as Long-Term Evolution (LTE/4G), the new fifth generation (5G) technology provides higher capacity, greater bandwidths, and lower latency (Niknejad et al., 2015). This new generation uses new sub 7 GHz frequency bands previously not used by LTE: 700 MHz and 3,6 GHz and higher frequencies like the 5G millimetre-wave (mm-Wave), referring to frequencies ranging from 24 GHz to 300 GHz (Rangan et al., 2014; Sim et al., 2020). In Europe, only the 26 GHz band is made available for international mobile telecommunication. These new frequencies deliver data faster but also result in smaller travel distances due to high signal attenuation (decrease of signal strength). To support these higher frequencies many new multiple output (MIMO) antennas, which use directional beamforming, are needed possibly changing the environmental RF-EMF exposure (Wali et al., 2022). With directional beamforming signals can be steered into a desired direction rather than traditional antennas which signals travel in all directions.

While the use of these 5G NR frequencies enables users to faster and more efficient mobile communication. They also raise concerns about increased environmental exposure, especially in densely populated areas. Understanding how these exposures vary under different environmental conditions, time and network providers is essential for the accurate risk assessment, public health research and informed regulatory decision making. Given the largely uncontrollable nature of RF-EMF exposure, international safety guidelines have been developed to define acceptable exposure limits.

Exposure guidelines for radiofrequency electromagnetic fields (RF-EMFs) are primarily established by the International Commission on Non-Ionizing Radiation Protection (ICNIRP) and the International Committee on Electromagnetic Safety (ICES) of the Institute of Electrical and Electronics Engineers (IEEE) (Ramirez-Vazquez et al., 2024). These international guidelines serve as the foundation for national regulatory standards worldwide. In the Netherlands, like many other European countries, the adopted exposure limits align with ICNIRP’s recommendations, specifying a maximum power flux density of 4.5 W/m² at 900 MHz and 9.0 W/m² at 1800 MHz for the general public (‘Guidelines for Limiting Exposure to Electromagnetic Fields (100 kHz to 300 GHz)’, 2020; Klimaat, 2021). These thresholds are designed to prevent established health effects such as tissue heating and are based on conservative assumptions to account for uncertainty in exposure response relationships.

Despite the existence of established exposure guidelines, research on real-world RF-EMF exposure under different environmental conditions remains inconclusive. This has created the need for more precise, scalable measurement methods. In response to this need, the ETAIN project has developed the ETAIN 5G Scientist app which uses the signal strength indicators (SSIs) from the phone, to estimate the RF-EMF downlink exposure. Downlink exposure refers to RF-EMF transmitted from antennas to the user’s mobile phone (Velghe et al., 2021). In contrast, uplink exposure refers to the signal transmitted from the phone to the antenna.

This app is a new approach to estimate and eventually map the environmental exposure levels of RF-EMF. Earlier research with this app demonstrated a positive log-linear relationship between Long Term Evolution (LTE) Signal Strength Indicators (SSI’s) and far-field RF-EMF exposure (Sandoval-Diez, 2024). Using calibration functions, these SSI’s can be used to estimate exposure originating from the phone. Next step is to perform more measurement campaigns where necessary to determine if the ETAIN 5G Scientist App could be a valuable tool for large scale exposure assessment. Evaluating the variability and uncertainty of app is essential to further assess this potential of this new approach.

Exposure measurement devices known as personal exposimeters (PEMs) are currently available for accurately measuring RF-EMF exposure (Röösli et al., 2010). One example is the ExpoM-RF 4, a frequency selective device specifically designed to quantify RF-EMF exposure across multiple frequency bands (ExpoM - RF 4 — Fields at Work, 2025). In this study, the ExpoM-RF 4 is used as a reference standard to validate the accuracy of the ETAIN 5G Scientist app’s exposure estimates.

Furthermore, environmental factors are known to influence RF-EMF exposure. Indoor environments often have more signal attenuation due to building structures than outdoor settings. Another article published in 2023 found that indoor downlink RF-EMF exposure varies randomly due to many uncontrollable factors as furniture, walls, doors and people moving causing reflection, diffractions, scatterings and signal attenuation. Emphasizing the high mobile and temporal variability unique to indoor environments (Mulugeta et al., 2023). RF-EMF exposure levels may fluctuate depending on the time of day. For example, peak usage during rush hours may increase signal traffic (Velghe et al., 2021). Thereby, exposure levels can vary depending on cell tower density. Urban areas typically have more coverage resulting in higher downlink exposure levels.

Additionally, differences in coverage between providers can lead to differences in estimated downlink exposure by the ETAIN 5G Scientist app. Since the app uses the mobile network associated with the provider’s SIM card, exposure levels can differ. Providers with weaker coverage may increase uplink exposure due to stronger transmission requirements from devices, while providers with strong coverage are likely to result in higher downlink exposure (Brzozek et al., 2021). In the Netherlands, three mobile network operators: KPN, Vodafone, and Odido operate their own national mobile infrastructures (Frequenties voor mobiele netwerken, 2022.; Markets, 2025). All three provide nationwide LTE and 5G coverage, though coverage quality and base station density vary by region, which in turn can influence the RF-EMF exposure measured by the ETAIN 5G Scientist app.

Investigating how these factors influence exposure is critical in evaluating the robustness of the ETAIN 5G Scientist app. This study aims to evaluate variability and uncertainty of app under real world conditions. Which results in the following main research question:

How does downlink radiofrequency electromagnetic field exposure variability differ across various conditions when measured with the ETAIN 5G Scientist app, and what is the uncertainty of the measurements made by the app compared to the ExpoM-RF 4?

Which is divided by the following Sub-Research Questions:

- How does the variability of downlink radiofrequency electromagnetic field exposure, measured using the ETAIN 5G Scientist app, differ across various conditions?

- Between different environmental conditions

  - Between rural and urban environments

  - Between indoor and outdoor environments

- Across time periods

- Between network providers (KPN, Vodafone, and Odido)

- What is the overall uncertainty in the downlink radiofrequency electromagnetic field exposure measurements made with the ETAIN Scientist 5G app compared to the ExpoM-RF 4

# Methods

## Study Design

This research is a field-based measurement study designed to assess real world variability and uncertainty of the ETAIN 5G Scientist app. The primary objective is to evaluate the reliability of this app in estimating downlink RF-EMF exposure under different environmental and temporal conditions. To evaluate the reliability of the ETAIN 5G Scientist App under realistic, real-world scenarios, measurements were conducted in different environmental and temporal conditions. These conditions were selected based on their known or expected influence on RF-EMF exposure levels, as documented in prior research. Variability between the individual conditions is compared for both the ETAIN 5G app and the ExpoM-RF 4 device.

Static and mobile measurements

Two types of measurements where conducted:

Static Measurements: Devices were placed on a fixed setup for a continuous period of 24 hours. These were performed across various combinations of urban and rural locations, Indoor and Outdoor setting and weekdays and weekends (Table 2.1).

Mobile Measurements: These measurements were carried out during a 15-minute short walk (7.5 minutes Outdoor followed by 7.5 minutes Indoor) around selected measurements locations. These were done in Urban and Rural locations (Table 2.2).

In total, 8 static and 4 mobile measurements were conducted across the Netherlands. In all the measurements three Motorola G54 5G smartphones (each equipped with a sim card of a different provider) and the ExpoM-RF 4 are measuring. Static measurement locations were selected based on personal connections, ensuring that the smartphones and ExpoM-RF 4 could be safely placed for 24 hours while also meeting the required environmental conditions. For mobile measurements, public locations in both urban and rural settings were chosen, where it was possible to walk both indoors and outdoors around the place.

Environmental and temporal conditions

To evaluate exposure variability under realistic conditions, measurements were in following different conditions: Indoor vs. Outdoor, fluctuations over 24 hours, network providers and urban versus rural settings. The last is based on the CBS urbanization classification using address density per km².

In this study, the degree of urbanization is based on address density numbers published by the “Central Bureau of Statistics” (CBS). This organization uses 5 categories in the classification of urbanization: Extremely urbanized, strongly urbanized, moderately urbanized, hardly urbanized and not urbanized (Degree of Urbanisation | CBS, 2025)

The address density per square kilometer is used to classify the urbanization degree in a specific neighborhood, district or municipality:

- Extremely urbanized: 2,500 addresses or more per square kilometer

- Strongly urbanized: 1,500 to 2,500 addresses per square kilometer

- Moderately urbanized: 1,000 to 1,500 addresses per square kilometer

- Hardly urbanized: 500 to 1,000 addresses per square kilometer

- Not urbanized: fewer than 500 addresses per square kilometer

Data referring to the addresses density in a neighborhood are from CBS district and neighborhood maps 2024 Version 1.(Statistiek, n.d.).

Measurements Overview

Table 2.1 RF-EMF Static Measurements overview in different environmental conditions

Table 2.2 RF-EMF Mobile Measurements overview in different environmental conditions

For further details on the documentation of each measurement, see Appendix A.1.

## Materials

Motorola G54 5G smartphone

Three Motorola G54 5G smartphones were used, each installed with the ETAIN 5G Scientist App. This phone is compatible with the 5G network. Each phone is connected to a different network provider (Vodafone, KPN and Odido) using a prepaid SIM card. Vodafone, KPN, and Odido are the only network providers in the Netherlands that operate their own mobile network infrastructure.

ETAIN 5G Scientist App
The ETAIN 5G Scientist App collects signal strength indicators, from both registered and neighboring cell towers. In these SSI’s two categories are specified: Reference Signal Received Power (RSRP)  and Received Signal Strength Indicator (RSSI) (Raida et al., 2018; Sandoval-Diez, 2024). These values are expressed in decibel milliwatts (dBm). RSRP values range from -140 to -43 dBm, and RSSI values range from -113 to -51 dBm.

RSSI reflect to the total received power and includes the reference signal, interference and thermal noise. RSRP is the received power of only the reference signal from the serving (or neighboring) cell. The serving cell, denoted with index 0, refers to the specific cell currently delivering the telecommunication signal. Neighboring cells, indexed from 1 till 9, are other nearby cell towers that the phone may switch to if the signal from the serving cell weakens. The serving and neighboring cells are all part of the same mobile network provider’s infrastructure.

In this study, only the RSRP values were used, as these include only the signal from the cell tower to the phone excluding other noise and interference. The ETAIN research group has developed two types of calibration formulas to convert the RSRP values into estimated downlink RF-EMF exposure expressed in power flux density (mW/m²) (Appendix B).

- The app’s calibration function is created to estimate the user’s downlink RF-EMF exposure which uses only the data from that specific device.

- The mapping calibration function uses measurements of multiple users, connected to different mobile data networks, to estimate the total downlink RF-EMF exposure at a given location. These estimations are then used to create exposure maps.

The measurement interval of the ETAIN 5G Scientist app may vary between phone models, but on the Motorola G54 5G smartphone, it records data every 5 seconds.

ExpoM-RF 4

The ExpoM-RF 4 measures in the frequency range of 50 MHz up to 6 GHz, allowing it to capture the majority of the bands used in the European mobile telecommunication system (ExpoM - RF 4 — Fields at Work, 2025). Measurements are recorded as Root Mean Square (RMS) values of the electric field strength and are expressed in volts per meter (V/m).

This device samples across all 35 predefined frequency bands, with each band being measured for 50 milliseconds during a continuous sampling interval. This interval can be changed on the utility program created by Fields at Work GmbH. The upper detection limit is fixed at 6 V/m and the lower detection limit depend on the specific frequency band but varies between 0.0009 V/m and 0.01 V/m.

Table 2.3 Materials for Static and Mobile measurements

| Item | Description |
| --- | --- |
| Three Motorola G54 5G phones | Measurement devices with the ETAIN 5G Scientist App |
| Three Prepaid Sim cards | Sim cards of Vodafone, Odido and KPN for each phone |
| ExpoM-RF 4 | PEM for accurate RF-EMF measurements |
| Wooden Table | Setup for static measurements |
| Ziplock bag | Protect the phones against rain during static measurements |
| Windows Laptop with the Fields at Work Utility program | Data retrieval from the ExpoM-RF 4 |
| Backpack | For carrying the measurement devices during mobile measurements |
| Notebook | Documentation of the static and mobile measurements |

## 2.3 Measurement Setup and Procedure

The measurement setup differed between static and mobile measurements. Static measurements were designed to capture long-term environmental exposure at a fixed location, whereas mobile measurements aimed to compare indoor and outdoor exposure within a short, controlled walking route. Both setups used a combination of three smartphones and the ExpoM-RF 4.

Static Measurements

For static measurements, the three smartphones and the ExpoM-RF 4 were placed on a wooden table measuring 42 cm in width, 30 cm in depth, and 52 cm in height (Table 2.3). The devices for all static measurements arranged as shown on the table.

Figure 2.3.1 Experimental setup for the static RF-EMF exposure measurements. Using a wooden table, three Motorola G54 5G phones and the ExpoM-RF 4.

Mobile Measurements

Mobile measurements were conducted while walking, using the same combination of three smartphones and the ExpoM-RF 4. All devices were placed inside a backpack, positioned as high and as far from the human body as possible to reduce body shielding by the researcher or bag material (Eeftens et al., 2020).

Measurement Procedure

The steps followed prior, during and after each individual measurement are presented in Appendix A.2. Different procedures were followed for the static and mobile measurements. These steps were consistently applied to ensure that each measurement was complete, securely stored, and in the correct format.

## 2.4 Data handling

The data analysis consists of several processing and conversion steps before variability and uncertainty can be assessed. Raw measurements from the ETAIN 5G Scientist app and the ExpoM-RF 4 device must first be converted into a standardized exposure metric: power flux density (mW/m²). The steps to do this are provided in Appendix B.1, the specific formulas used in Appendix B.2. The measurement of 5G frequencies by the ETAIN 5G Scientist App is very limited in the Netherlands, as phones are mostly connected to LTE network bands. For this reason, the conversion steps have only been performed with the LTE measurements from the app.

After all data has been converted and expressed in a consistent exposure metric (power flux density in mW/m²), preparation for the analysis can begin. Each measurement is filtered so that the dataset only has measurements between start and end time (see Appendix A1). Incorporate the measurement condition to each data set: rural or urban settings, location name and indoor or outdoor location (only for static measurements).

Static Measurements

In the analysis of the static measurements the 24-hour measurements are separated by creating time frames of 3 hours during the day and one time frame during the night: 06:30-09:30, 09:30-12:30, 12:30-15:30, 15:30-18:30, 18:30-21:30, 21:30-6:30. Time frames 09:30-12:30 and 18:30-21:30 are in this analysis defined as rush hours.

Mobile Measurements

In the mobile measurements, the data is divided into two groups: Indoor and Outdoor measurements. This will be done by using the timestamp of the measurements data. The first 7.5 minutes are outdoor and the second 7.5. minutes are indoor.

Mobile measurement 4 (Spar Cothen) was excluded from the uncertainty analysis due to the loss of the timestamp data. The definition of the seconds was erased, making 10 seconds timeframes and alignment with the ExpoM-RF 4 data not possible which is needed for the mapping calibration steps (Appendix B1.2).

## 2.5 Variability analysis

Variability across conditions is assessed by comparing estimated downlink RF-EMF exposure, calibrated with the App Calibration Function (see Appendix B1.1), with the ExpoM-RF 4. These comparisons reveal trends across environmental and temporal conditions, and provider.

Variability within one measurement location

Start by calculating statistical parameters for each group, distinguishing between time frames and providers for static measurements, and between indoor and outdoor conditions for mobile measurements. Perform these calculations separately for both the ExpoM-RF 4 and the ETAIN 5G Scientist app. Since the measurements cannot be assumed to follow a normal distribution, nonparametric statistics are used: specifically, the median, median absolute deviation (MAD), and interquartile range (IQR).

For static measurements, analyze variability by comparing trends across different conditions (time frames and providers), and compare these trends between the ETAIN 5G Scientist app and the ExpoM-RF 4. For mobile measurements, assess differences between indoor and outdoor environments, as well as between providers, and again compare these trends with those observed using the ExpoM-RF 4. Trends are visualized using boxplots. When comparing providers, a Wilcoxon signed-rank test is performed to determine whether the exposure estimates differ significantly. Trends are visualized using boxplots. When comparing providers, a paired Wilcoxon signed-rank test is performed to determine whether the exposure estimates differ significantly.

Variability between all locations

When comparing multiple static and mobile measurements, all data are combined to analyze trends across various conditions, as well as between the ExpoM-RF 4 and the ETAIN 5G Scientist app. The goal is to visualize patterns across conditions in both static and mobile measurements.

For static measurements, comparisons are made between rural and urban settings, indoor and outdoor locations, different time frames during the day and night, and between providers. These comparisons are visualized using boxplots, and statistical significance between conditions is assessed using the paired Wilcoxon signed-rank test. For mobile measurements, comparisons are made between rural and urban environments, indoor and outdoor locations, and across providers. These trends are then compared with those from the ExpoM-RF 4 to evaluate whether the variability and patterns are consistent between the two devices.

## 2.6 Uncertainty analysis

To evaluate uncertainty, the Mapping Calibration Function is used to derive total environmental downlink exposure from the app using data from all three network providers, which is then compared to reference measurements from the ExpoM-RF 4 (see appendix B1.2). This comparison quantifies the measurement uncertainty of the app.

Uncertainty within one measurement location

To assess uncertainty within individual measurement locations, a direct comparison is made between the ETAIN 5G Scientist app and the ExpoM-RF 4. For this analysis, the mapping calibration is applied, and measurements from all three providers are combined. At each measurement location, the differences between the app and the reference device are visualized across time frames for static measurements and across indoor and outdoor conditions for mobile measurements. A paired Wilcoxon signed-rank test is then conducted to evaluate whether these differences are statistically significant. Finally, a correlation plot is created to assess whether a consistent relationship exists between the two measurement methods at that location.

Uncertainty between all locations

For the overall comparison, all measurements are combined and annotated with their corresponding conditions. Differences between conditions are then visualized. For static measurements, boxplots are created to compare urban vs. rural and indoor vs. outdoor conditions within a single plot. This is followed by additional boxplots that show these differences over time. For mobile measurements, a single boxplot is used to compare indoor, outdoor, urban, and rural conditions, and these are compared against the ExpoM-RF 4. Within each condition, a paired Wilcoxon signed rank test is used to see if the differences are significant.

In addition, correlation plots are generated to evaluate the strength of the relationship between the app’s measurements and those from the ExpoM-RF 4. To further assess agreement between the two devices, a Bland–Altman plot is created. This plot calculates the mean difference between paired measurements (Dewitte et al., 2002), with the x-axis representing the average of the two measurements and the y-axis showing the difference between them. The Bland–Altman plot provides insight into the average error, reveals any systematic bias in the app’s measurements, and indicates whether accuracy varies depending on exposure levels.

The full statistical analysis was performed using R (version 4.5.1) within Rstudio (version: 2025.05.1+513, released: 2025-06-05)

# Results

This section presents the variability and uncertainty in estimated downlink RF-EMF exposure as measured with the ETAIN 5G Scientist App and validated using ExpoM-RF 4.

## 3.1 Variability Static measurements

Variability in each location

For each of the eight static measurements (see Table 2.1.1 for locations and conditions), individual analyses were performed to compare exposure patterns over time and across providers between app and ExpoM-RF 4. Exposure levels recorded by the ETAIN 5G Scientist app appear to fluctuate over time within a single location, whereas the ExpoM-RF 4 shows relatively stable measurements with low variance. Additionally, large significant differences between providers are observed and consistently visualized across all measurement locations. The results are presented in Appendix C.1, boxplots and summary tables for each measurement location provide insight into within location variability.

Comparison between Indoor and Outdoor locations

Overall, Figures 3.1.1 and 3.1.2 show the downlink RF-EMF exposure levels measured respectively with the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations, separated by provider and urban and rural settings.

For both figures, exposure levels were significantly higher in urban settings across all providers and environments (p < 0.0001). This trend is most pronounced in outdoor locations and for the providers KPN and Vodafone. Odido had very low estimations in all environmental conditions, but the same trend is still visible. Specifically, Figure 3.1.1 shows higher variability in exposure in urban outdoor areas, with KPN showing the highest median exposure values. Figure 3.1.2 confirms the same trend, though estimations are lower. Measurements in indoor locations where constructively low in both ExpoM-RF 4 and ETAIN app, especially in rural locations. For full descriptive statistics including medians, MAD and IQR, see Appendix E, Tables E1.1 and E1.2.

Figure 3.1.1 Comparison of estimated downlink RF-EMF exposure (mW/m²) across different environments (indoor vs. outdoor), settings (rural vs. urban), and mobile network providers (KPN, Odido, Vodafone) using the ETAIN 5G Scientist App. Statistical significance between rural and urban is indicated with: **** = p < 0.0001. For descriptive statistics see appendix E, Table E1.1.

Figure 3.1.2 Comparison of downlink RF-EMF exposure (mW/m²) across different environments (indoor vs. outdoor) and settings (rural vs. urban) using the ExpoM-RF 4. Statistical significance between rural and urban is indicated with: **** = p < 0.0001. For descriptive statistics see appendix E, Table E1.2.

Comparison between temporal differences
Figures 3.1.3 and 3.1.4 show the downlink RF-EMF exposure levels measured respectively with the ETAIN 5G Scientist App and the ExpoM-RF 4 plotted over time, separated by provider and grouped by urban and rural settings.

Urban settings show higher exposure levels than rural across all providers, which is again in line with the ExpoM-RF 4. Temporal patterns, also show lower exposure levels during the night compared to daytime. Though differences between providers during daytime hours appear limited in both ExpoM-RF 4 and app’s data. For full descriptive statistics, including medians, MAD, and IQR values, see Appendix E, Tables E1.3 and E1.4.

Figure 3.1.3 Comparison of estimated downlink RF-EMF exposure (mW/m²) across timeframes and settings (rural vs. urban) for each mobile network provider (KPN, Odido, Vodafone) using the ETAIN 5G Scientist App. Statistical significance between rural and urban is indicated with: * = p < 0.05, *** =  p < 0.001 and **** = p < 0.0001. For descriptive statistics see appendix E, Table E1.3.

Figure 3.1.4 Comparison of downlink RF-EMF exposure (mW/m²) across different timeframes and settings (rural vs. urban) using the ExpoM-RF 4. Statistical significance between rural and urban is indicated with **** = p < 0.0001. For descriptive statistics see appendix E, Table E1.4.

Comparison between providers

Figure 3.1.5 presents a comparison of estimated downlink RF-EMF exposure levels across the three major mobile network providers, KPN, Odido, and Vodafone separately for indoor and outdoor locations, using the ETAIN 5G Scientist App. Exposure levels differed significantly between providers in both indoor and outdoor locations (p < 0.0001). The highest measured exposure levels were observed for KPN, followed by Vodafone and then Odido.

Figure 3.1.5 Comparison of estimated downlink RF-EMF exposure (mW/m²) across mobile network provider (KPN, Odido, Vodafone) and environments (indoor versus outdoor) for each using the ETAIN 5G Scientist App. Statistical significance between rural and urban is indicated with: **** = p < 0.0001. For descriptive statistics see appendix E, Table E1.1.

## 3.2 Variability Mobile measurements

Each location
For each of the four mobile measurements (see Table 2.1.2 for locations and conditions), individual analyses were performed to compare exposure patterns over time and across providers between app and ExpoM-RF 4. RF-EMF exposure levels in individual mobile measurements appear to be more stable overall. A clear trend is observed between indoor and outdoor measurements, which is also reflected in the data from the ExpoM-RF 4. When comparing providers in indoor locations, the measurements tend to align more closely, often presenting no significant differences between providers. In contrast, outdoor measurements show greater variability between providers, with more pronounced and statistically significant differences. The results are presented in Appendix C.2, boxplots and summary tables for each measurement location provide insight into each location’s variability.

Comparison between Indoor and Outdoor locations
Figures 3.2.1 and 3.2.2 show the downlink RF-EMF exposure measured during mobile measurements with respectively the ETAIN 5G Scientist App and the ExpoM-RF 4, across indoor and outdoor locations, separated by rural and urban settings and by provider.

In Figure 3.2.1, the ETAIN 5G App estimates reveal significantly higher exposure levels in urban outdoor environments for Vodafone (p < 0.0001), while differences are not statistically significant for KPN and Odido. Indoor exposures remained consistently low across all providers and settings differences between urban and rural were only significant for KPN. While the spread of exposure values increases in urban measurements, particularly outdoors, the median values between rural and urban locations remain relatively close.

Figure 3.2.2 confirms this overall trend when measured by the reference device, with no significant differences between urban and rural in indoor and outdoor locations. Furthermore, exposure levels and variability are visibly higher in urban outdoor areas, these measurements from the ExpoM-RF 4 are in line with the app’s results.

For full descriptive statistics, including medians, MAD, and IQR values, see Appendix E, Tables E2.1 and E2.2.

Figure 3.2.1 Comparison of estimated downlink RF-EMF exposure (mW/m²) across different environments (indoor vs. outdoor), settings (rural vs. urban), and mobile network providers (KPN, Odido, Vodafone) using the ETAIN 5G Scientist App. Statistical significance between rural and urban is indicated with **** = p < 0.0001, ‘ns’ indicates no significant differences. For descriptive statistics see appendix E, Table E2.1.

Figure 3.2.2 Comparison of downlink RF-EMF exposure (mW/m²) across different environments (indoor vs. outdoor) and settings (rural vs. urban) using the ExpoM-RF 4. ‘ns’ indicates no significant differences between rural and urban settings. For descriptive statistics see appendix E, Table E2.2.

Comparison between providers
Figure 3.2.3 presents a comparison of estimated downlink RF-EMF exposure levels across the three major mobile network providers, KPN, Odido, and Vodafone, for indoor and outdoor locations.

In the indoor locations, exposure levels were significantly different between all providers (p < 0.0001). In the outdoor locations, exposure levels were overall higher. Vodafone showed significantly different exposure levels (p < 0.0001) in comparison with KPN and Odido, while no significant difference was observed between KPN and Odido (ns).

Figure 3.2.3 Comparison of estimated downlink RF-EMF exposure (mW/m²) across mobile network provider (KPN, Odido, Vodafone) and environments (indoor versus outdoor) for each using the ETAIN 5G Scientist App. Statistical significance between rural and urban is indicated with: **** = p < 0.0001 and ‘ns’ indicates no significant differences. For descriptive statistics see appendix E, Table E2.1.

## 3.3 Uncertainty Static measurements

Uncertainty in each location

For each of the eight static measurements, an individual analyses were performed to compare exposure levels over time between App and ExpoM-RF 4. At each measurement location, the ETAIN 5G Scientist app appears to overestimate downlink RF-EMF exposure compared to the ExpoM-RF 4. In all individual static measurements, large and statistically significant differences are observed. Moreover, when assessing the relationship between the ExpoM-RF 4 and the ETAIN 5G Scientist app, little to no correlation is found, often resulting in negative or near-zero R values. The results are presented in Appendix D.1, boxplots, correlation plots and summary tables for each measurement location provide insight into within location uncertainty.

Uncertainty analysis all locations
Figure 3.3.1 compares total downlink RF-EMF exposure levels (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across different locations (indoor vs outdoor) and settings (rural vs urban) during static measurements. Figure 3.3.2 presents the same data distributed over six timeframes, separately for rural and urban settings.

Across all conditions, the app consistently estimated higher exposure levels than the ExpoM-RF 4, with statistically significant differences in each group (p < 0.0001). The greatest variability in the app’s measurements was observed in urban outdoor environments, while the highest exposure levels were found in urban indoor environments. In contrast, ExpoM-RF 4 measurements remained low and stable across all environments, with slightly elevated values in urban outdoor conditions.

Figure 3.3.2 further illustrates that exposure levels measured by the ETAIN App are consistently higher in urban areas but show limited variation across timeframes. ExpoM-RF 4 values also reflect this trend but show a small decrease in exposure during nighttime.

For full descriptive statistics, including medians, MAD, and IQR values, see Appendix E, Tables E3.1 and E3.2.

Figure 3.3.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) between the ETAIN 5G Scientist App and the ExpoM-RF 4 across different environments (indoor vs. outdoor) and settings (rural vs. urban). Statistical significance between measurement sources is indicated with: **** = p < 0.0001. For descriptive statistics see appendix E, Table E3.1.

Figure 3.3.2 Comparison of total environmental downlink RF-EMF exposure (mW/m²) between the ETAIN 5G Scientist App and the ExpoM-RF 4 across different timeframes and settings (rural vs. urban). Statistical significance between measurement sources is indicated with: **** = p < 0.0001. For descriptive statistics see appendix E, Table E3.2.

Correlation and Agreement

Figure 3.3.3 presents a scatter plot with a regression line comparing static downlink RF-EMF exposure estimates from the ETAIN 5G Scientist App with measurements from the ExpoM-RF 4. The correlation coefficient is R = 0.76 (p < 2.2e−16), indicating a moderate to strong positive correlation.

Figure 3.3.4 shows a Bland–Altman plot, which assesses agreement between the two measurement methods. The x-axis displays the mean exposure of each paired measurement, while the y-axis shows the difference between the ExpoM-RF 4 and the app. The blue dashed line indicates the median difference, and the red dashed lines represent the 95% limits of agreement. The mean difference is consistently below zero. Furthermore, as the mean exposure increases, the difference between the app and the ExpoM-RF 4 also decreases.

Figure 3.3.3 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 during static measurements. A significant strong positive correlation was observed (R = 0.76, p < 2.2e−16).

Figure 3.3.4 Bland–Altman plot comparing total downlink RF-EMF exposure (mW/m²) between the ETAIN 5G Scientist App and the ExpoM-RF 4 during static measurements. The x-axis shows the mean exposure, while the y-axis shows the difference between the two measurement devices. The blue dashed line indicates the mean difference, red dashed lines represent the 95% limits of agreement.

## 3.4 Uncertainty Mobile measurements

Uncertainty in each location
For each of the three mobile measurements, an individual analyses were performed to compare exposure levels in indoor and outdoor locations between App and ExpoM-RF 4. At all individual locations, indoor measurements showed no statistically significant differences between the ETAIN 5G Scientist app and the ExpoM-RF 4. In contrast, outdoor measurements did reveal significant differences between the two devices. When assessing correlation within individual locations, a clearer pattern emerges, often indicating a moderate correlation between the app and the reference device.The results are presented in Appendix D.2, boxplots, correlation plots and summary tables for each measurement location provide insight into within location uncertainty.

Uncertainty analysis all locations
Figure 3.4.1 compares total downlink RF-EMF exposure (mW/m²) as estimated by the ETAIN 5G Scientist App and measured by the ExpoM-RF 4 across rural and urban settings, for both indoor and outdoor locations during mobile measurements.

Urban outdoor exposure levels measured by the ExpoM-RF 4 have greater spread than those estimated by the app. However, these differences were not significant. In indoor and rural settings, no significant differences were observed between the two methods (ns). Outdoor rural and indoor urban settings displayed a significant difference between ExpoM-RF 4 and app.

For full descriptive statistics, including medians, MAD, and IQR values, see Appendix E, Tables E3.4.

Figure 3.4.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) between the ETAIN 5G Scientist App and the ExpoM-RF 4 across different environments (indoor vs. outdoor) and settings (rural vs. urban). Statistical significance between measurement sources is indicated with: *** = p < 0.001, * = p < 0.05 and ‘ns’ = not significant. For descriptive statistics see appendix E, Table E3.4.

Correlation and Agreement

Figure 3.4.2 presents a scatter plot with a regression line comparing mobile environmental downlink RF-EMF exposure estimates from the ETAIN 5G Scientist App with measurements from the ExpoM-RF 4. The correlation coefficient is R = 0.51 (p < 2.2e−16), indicating a moderate positive correlation.

Figure 3.4.3 shows a Bland–Altman plot assessing the agreement between the ETAIN App and the ExpoM-RF 4 in mobile measurements. The mean of each paired exposure measurement is plotted on the x-axis, and the difference between ExpoM-RF 4 and the ETAIN App is shown on the y-axis. The blue dashed line represents the median difference, and the red dashed lines indicate the 95% limits of agreement. Furthermore, as the mean exposure increases, the difference between the app and the ExpoM-RF 4 also increases.

Figure 3.4.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 during mobile measurements. A significant moderate positive correlation was observed (R = 0.51, p < 2.2e−16).

Figure 3.4.3 Bland–Altman plot comparing total downlink RF-EMF exposure (mW/m²) between the ETAIN 5G Scientist App and the ExpoM-RF 4 during mobile measurements. The x-axis shows the mean exposure, while the y-axis shows the difference between the two measurement devices. The blue dashed line indicates the mean difference, red dashed lines represent the 95% limits of agreement.

# Conclusion and Discussion

This study aimed to evaluate variability and uncertainty in downlink RF-EMF exposure using the ETAIN 5G Scientist App and the ExpoM-RF 4 across diverse real-world conditions. Outdoor locations showed increased exposure levels in both the ExpoM-RF 4 and the app compared to indoor environments. These findings were clearly observed across all static measurements and were also evident in the mobile measurements when comparing all conditions collectively. Additionally, these differences were consistently present when analyzing exposure levels at each individual measurement location, for both the reference device and the app. When comparing exposure levels in each mobile measurement, all comparisons were significant. Outdoor environments displayed greater variation in app and ExpoM-RF 4 measurements than indoor. But this difference is greater in measurements done by the ETAIN 5G Scientist app. These results align with the expectations, as downlink RF-EMF signals are unable to penetrate building exteriors effectively, resulting in signal attenuation and lower exposure levels indoors. The difference in variability between the ETAIN 5G Scientist App and the ExpoM-RF 4 can be due to the app relying on a single provider’s network which is more prone to changes in connection with the antennas.

In urban and rural settings, the results are largely in line with the comparison between indoor and outdoor locations. Measurements in urban areas were always higher than in rural, in both mobile and static measurements. This shows the app follows the same trends as the ExpoM-RF 4. These findings are likely due to higher cell tower density increasing downlink RF-EMF exposure levels. A greater variation in urban than in rural locations was observed in both the app and the ExpoM-RF 4. The app’s measurements showed greater variation compared to those from the ExpoM-RF 4 in both urban and rural settings.

RF-EMF exposure levels can vary significantly over the course of the day and night, influenced by changes in user behavior, network demand, and environmental conditions. In this study, temporal differences during the day (06:30 to 21:30) were higher than during the night in both app and ExpoM-RF 4, likely due to decreased mobile network usage at night. Differences between specific timeframes throughout the day were harder to define. However, exposure levels tended to increase throughout the day, peaking between 15:30 and 18:30 in both urban and rural settings according to the ExpoM-RF 4 measurements. These patterns were not clearly observed in the ETAIN 5G Scientist App data, suggesting that the app is less reliable in capturing temporal variations during the day. The spread of measurements from the app was also considerably larger compared to the ExpoM-RF 4, which aligns with variability results in different environmental conditions. Additionally, measurement variation increased during rush hours compared to non-rush hours, indicating that the app’s estimations become more unstable when mobile network demand increases.

When comparing providers, substantial differences in exposure were observed. All pairwise comparisons showed statistically significant differences, except for the mobile measurements between KPN and Odido in outdoor locations. In the static measurements, KPN exhibited the highest exposure levels and variability, followed by Vodafone and then Odido. In the mobile measurements, KPN again showed the highest exposure in outdoor settings, followed by Odido and Vodafone. Interestingly, in indoor locations, Odido had the highest exposure levels.

When estimating total environmental exposure, large and significant differences were observed. In the static measurements, the ETAIN 5G Scientist App consistently overestimated exposure levels compared to the ExpoM-RF 4 across all conditions. In many cases, the app’s estimations were approximately ten times higher. While the overall correlation between the two devices was moderate to strong, this pattern did not consistently hold across individual locations. In fact, correlation coefficients at specific locations were often below 0.5 or even negative. The large sample size across diverse conditions reveals a general trend between the ExpoM-RF 4 and the app, but this trend does not consistently emerge within individual static locations. The Bland–Altman plot confirms the systematic overestimation by the ETAIN 5G Scientist App. At higher exposure levels, the app’s estimations become increasingly inaccurate, which is consistent with earlier findings.

In the mobile measurements, the app appeared to underestimate exposure levels compared to the reference device. However, this trend was primarily driven by a single mobile measurement (see Appendix D, Figure D2.2.1), where many students were actively using the mobile network. This likely resulted in elevated downlink RF-EMF exposure levels captured by the ExpoM-RF 4, which were difficult for the app to measure, as it was operating with only three phones. Differences between the app and the ExpoM-RF 4 were not significant across urban outdoor and rural indoor locations, and median values appeared to be at a similar level in both devices (Figure 3.4.1, Table (E3.4). The overall correlation between the app and the ExpoM-RF 4 was moderate, though stronger correlations were observed in individual mobile locations for example. The Bland–Altman plot for mobile data also confirms that estimation accuracy decreases as exposure levels increase.

The ETAIN 5G Scientist App demonstrates considerable potential as a tool for public engagement and large-scale RF-EMF exposure monitoring. By leveraging the widespread availability of smartphones, the app supports a citizen science approach to data collection that is both cost-effective and scalable. As nearly everyone carries a mobile phone, this method allows for high-resolution data to be collected across temporal and spatial scales with minimal additional effort. Unlike traditional measurement campaigns that rely on costly exposimeters, like the ExpoM-RF 4, the app enables large groups of participants to contribute exposure data passively, while also receiving personalized feedback on their own exposure levels. This dual function enhances user awareness and promotes engagement with the topic of enviromental RF-EMF exposure.

A key strength of the current study lies in its use of real-world conditions to evaluate the app’s performance. By including both static and mobile measurements, the study captured a broad range of RF-EMF exposure scenarios. Static 24-hour measurements allowed for the observation of temporal variation across day and night cycles, while mobile measurements reflected typical user behavior and app usage patterns. The combination of these two approaches provided a comprehensive picture of how the app performs in practical contexts. Furthermore, the potential to aggregate large volumes of smartphone-based data opens opportunities for identifying spatial exposure patterns or localized hotspots, offering insights that would be difficult to achieve through traditional measurement methods.

However, this study also had some limitations. The GPS coordinates used for the mobile measurements to distinguish between indoor and outdoor locations had an inaccuracy that could vary up to 100 meters. Due to this issue, only the timestamp was used to differentiate between indoor and outdoor measurements. When comparing temporal differences, timeframes of three hours were created during the day. As a result, the timeframes in this study did not fully align with those used in previous studies, potentially leading to missed patterns. Although the ExpoM-RF 4 was used as a reference device and can measure RF-EMF accurately, its measurements can still be influenced by body shielding (in mobile measurements) and signal reflection.

When reviewing prior literature, researcher reported lower downlink exposure in indoor environments compared to outdoor settings, supporting the outcomes of this study (Sagar et al., 2018). Another study published in 2016 challenged the common assumption that urban areas consistently exhibit higher RF-EMF exposure levels due to greater cell tower density (Martens et al., 2016). Measurements in rural locations showed greater exposure in less urbanized settings, possibly due to less attenuation and better signal propagation. Conversely, a more recent multi-country study using an RF-EMF sensor found that downlink signal power was indeed higher in large cities than in villages, aligning with the expected trend and the findings of this study. This is likely due to denser antenna deployment in urban environments (Bladel et al., 2025). These studies stated that the contrasting findings highlight the complexity of RF-EMF exposure dynamics, which depend not only on cell tower density, but also on local conditions such as physical surroundings, building obstructions, and line of sight to antennas. In this study, the higher antenna density in both ExpoM-RF 4 and ETAIN 5G Scientist app data is likely the reason for the increased downlink RF-EMF exposure levels in urban settings.

Several studies have found that exposure is typically higher during daytime hours, especially during peak usage periods such as morning (06:30–09:30) and evening (16:30–19:30) rush hours, when network traffic increases (Velghe et al., 2021). Similar reports from a study done in 2009, concluding that daytime exposure was twice as high as nighttime exposure. With mobile phones contributing more during the day and base stations dominating at night (Frei et al., 2009). Another study observed an overall increase in outdoor downlink exposure over time, and confirmed that daytime exposure was generally 10 to 30% higher than at night, citing growing infrastructure and traffic patterns as key factors (Urbinello et al., 2014). The results in these previous are in line with the results of this study specifically the measurements performed by the ExpoM-RF 4.

Recent benchmarking from the NL Mobile Network Test Report (2024) indicates that KPN leads in coverage reach (98.4%), followed closely by Odido (97.7%) and Vodafone (97.0%) (Netherlands, March 2022, Mobile Network Experience Report | Opensignal, 2022.; The 2024 Mobile Network Test in the Netherlands, 2025). These metrics are relevant for downlink RF-EMF exposure, as networks with broader and denser infrastructure can result in higher environmental exposure. When comparing these benchmarks to the results of this study, a possible relationship emerges between network coverage reach and measured exposure levels KPN not only had the highest coverage but also the highest measured exposure, possibly explaining these differences although the differences in coverage reach are relatively small.

One possible explanation to the large difference in static measurements relates to the phone’s usage state during the static measurements, the phones remained in idle or non-use mode, whereas the calibration functions were originally developed under active usage condition, for example, with phones running applications like Google Maps or streaming videos. In the mobile measurements, this discrepancy was less pronounced. The app tended to underestimate exposure in environments with high user activity and fluctuating network demand, but in rural areas, measurements were more closely aligned with those from the ExpoM-RF 4. These findings suggest that the app’s calibration performs more reliably under stable, low-to-moderate exposure conditions, but struggles in real-world scenarios with high exposure levels.

These findings highlight the need to refine the calibration functions. Although the original calibration formulas were developed using data from the Netherlands, Switzerland, Spain, Greece, and Germany, slight differences between countries were observed during this study (Sandoval-Diez, 2024). Variations in network infrastructure, such as cell tower density, geographical layout, and country-specific band usage, may contribute to differences between the ExpoM-RF 4 and the app’s measurements.

In future research, a cell tower density analysis could be conducted to investigate whether a relationship exists between local cell tower density and the exposure levels of the three providers. Additionally, this study only used RSRP values, a full analysis could also incorporate RSSI values to determine whether they result in more accurate exposure estimations. Eventually, 5G NR (New Radio) values can also be included and compared with ExpoM-RF 4 measurements to assess differences in exposure between 5G and LTE networks. Lastly, this dataset could be used to develop an improved version of the app, including an updated mapping calibration function specifically adapted to the Dutch mobile network environment.

# GenAI disclosure statement

In the preparation and writing of this report, I used ChatGPT to support the following:

- Improving writing clarity, grammar, punctuation and fluency

- Generating synonyms and rephrasing examples for enhanced readability

- Providing short explanations of technical terms or principles for my own knowledge

- Brainstorming and refining ideas for titles, research questions, and document structure

- Receiving assistance with R code when necessary

All content was critically reviewed by me as the author. The AI tool functioned solely as a supportive instrument.

# References

- 5G Technology. (n.d.). Retrieved 26 May 2025, from https://map.sciencemediahub.eu/5g

- Antenneregister. (n.d.). Retrieved 16 October 2024, from https://antenneregister.nl/Html5Viewer/Index.html?viewer=Antenneregister%5Fextern

- Beekhuizen, J., Heuvelink, G. B. M., Huss, A., Bürgi, A., Kromhout, H., & Vermeulen, R. (2014). Impact of input data uncertainty on environmental exposure assessment models: A case study for electromagnetic field modelling from mobile phone base stations. Environmental Research, 135, 148–155. https://doi.org/10.1016/j.envres.2014.05.038

- Bhatt, C. R., Thielens, A., Billah, B., Redmayne, M., Abramson, M. J., Sim, M. R., Vermeulen, R., Martens, L., Joseph, W., & Benke, G. (2016). Assessment of personal exposure from radiofrequency-electromagnetic fields in Australia and Belgium using on-body calibrated exposimeters. Environmental Research, 151, 547–563. https://doi.org/10.1016/j.envres.2016.08.022

- Bladel, H. V., Stroobandt, B., Veludo, A. F., Deprez, K., Röösli, M., Tognola, G., Parazzini, M., Thuróczy, G., Polańska, K., Politański, P., Wiart, J., Guxens, M., & Joseph, W. (2025). RF-EMF exposure assessment with add-on uplink exposure sensor in different microenvironments in seven European countries. Environment International, 197, 109368. https://doi.org/10.1016/j.envint.2025.109368

- Brzozek, C., Zeleke, B. M., Abramson, M. J., Benke, K. K., & Benke, G. (2021). Radiofrequency electromagnetic field exposure assessment: A pilot study on mobile phone signal strength and transmitted power levels. Journal of Exposure Science & Environmental Epidemiology, 31(1), 62–69. https://doi.org/10.1038/s41370-019-0178-6

- Chiaramello, E., Bonato, M., Fiocchi, S., Tognola, G., Parazzini, M., Ravazzani, P., & Wiart, J. (2019). Radio Frequency Electromagnetic Fields Exposure Assessment in Indoor Environments: A Review. International Journal of Environmental Research and Public Health, 16(6), Article 6. https://doi.org/10.3390/ijerph16060955

- Degree of urbanisation | CBS. (2025). https://www.cbs.nl/en-gb/our-services/methods/definitions/degree-of-urbanisation

- Dewitte, K., Fierens, C., Stöckl, D., & Thienpont, L. M. (2002). Application of the Bland–Altman Plot for Interpretation of Method-Comparison Studies: A Critical Investigation of Its Practice. Clinical Chemistry, 48(5), 799–801. https://doi.org/10.1093/clinchem/48.5.799

- Eeftens, M., Dongus, S., Bürgler, A., & Röösli, M. (2020). A real-world quality assessment study in six ExpoM-RF measurement devices. Environmental Research, 182, 109049. https://doi.org/10.1016/j.envres.2019.109049

- ExpoM - RF 4—Fields at Work. (2025). https://fieldsatwork.ch/products/expom-rf/expom-rf4/

- Frei, P., Mohler, E., Neubauer, G., Theis, G., Bürgi, A., Fröhlich, J., Braun-Fahrländer, C., Bolte, J., Egger, M., & Röösli, M. (2009). Temporal and spatial variability of personal exposure to radio frequency electromagnetic fields. Environmental Research, 109(6), 779–785. https://doi.org/10.1016/j.envres.2009.04.015

- Frequenties voor mobiele netwerken. (n.d.). Antennekaart.nl. Retrieved 26 May 2025, from https://antennekaart.nl/page/frequencies

- Guidelines for Limiting Exposure to Electromagnetic Fields (100 kHz to 300 GHz). (2020). Health Physics, 118(5), 483–524. https://doi.org/10.1097/HP.0000000000001210

- Huss, A., Dongus, S., Aminzadeh, R., Thielens, A., van den Bossche, M., Van Torre, P., de Seze, R., Cardis, E., Eeftens, M., Joseph, W., Vermeulen, R., & Röösli, M. (2021). Exposure to radiofrequency electromagnetic fields: Comparison of exposimeters with a novel body-worn distributed meter. Environment International, 156, 106711. https://doi.org/10.1016/j.envint.2021.106711

- Klimaat, M. van E. Z. en. (2021, March 30). Blootstellingslimieten voor elektromagnetische velden—Straling en gezondheid—Antennebureau [Webpagina]. Ministerie van Economische Zaken en Klimaat. https://www.antennebureau.nl/straling-en-gezondheid/blootstellingslimieten-voor-elektromagnetische-velden

- Korkmaz, E., Aerts, S., Coesoij, R., Bhatt, C. R., Velghe, M., Colussi, L., Land, D., Petroulakis, N., Spirito, M., & Bolte, J. (2024). A comprehensive review of 5G NR RF-EMF exposure assessment technologies: Fundamentals, advancements, challenges, niches, and implications. Environmental Research, 260, 119524. https://doi.org/10.1016/j.envres.2024.119524

- Markets, R. and. (2025, March 28). Netherlands Telecom (Fixed Telephony, Broadband, Mobile, Pay-TV) Operators Intelligence Report 2025, Competitive Analysis of KPN, Odido Netherlands, VodafoneZiggo, lycamobile, Canal Digitaal, Lebara. GlobeNewswire News Room. https://www.globenewswire.com/news-release/2025/03/28/3051273/28124/en/Netherlands-Telecom-Fixed-Telephony-Broadband-Mobile-Pay-TV-Operators-Intelligence-Report-2025-Competitive-Analysis-of-KPN-Odido-Netherlands-VodafoneZiggo-lycamobile-Canal-Digitaal.html

- Martens, A. L., Slottje, P., Meima, M. Y., Beekhuizen, J., Timmermans, D., Kromhout, H., Smid, T., & Vermeulen, R. C. H. (2016). Residential exposure to RF-EMF from mobile phone base stations: Model predictions versus personal and home measurements. Science of The Total Environment, 550, 987–993. https://doi.org/10.1016/j.scitotenv.2016.01.194

- Mulugeta, B. A., Wang, S., Ben Chikha, W., Liu, J., Roblin, C., & Wiart, J. (2023). Statistical Characterization and Modeling of Indoor RF-EMF Down-Link Exposure. Sensors, 23(7), Article 7. https://doi.org/10.3390/s23073583

- Netherlands, March 2022, Mobile Network Experience Report | Opensignal. (n.d.). Retrieved 27 May 2025, from https://www.opensignal.com/reports/2022/03/netherlands/mobile-network-experience

- Niknejad, A. M., Thyagarajan, S., Alon, E., Wang, Y., & Hull, C. (2015). A circuit designer’s guide to 5G mm-wave. 2015 IEEE Custom Integrated Circuits Conference (CICC), 1–8. https://doi.org/10.1109/CICC.2015.7338410

- R Studio. (2025). Posit. https://www.posit.co/

- Raida, V., Lerch, M., Svoboda, P., & Rupp, M. (2018). Deriving Cell Load from RSRQ Measurements. 2018 Network Traffic Measurement and Analysis Conference (TMA), 1–6. https://doi.org/10.23919/TMA.2018.8506494

- Ramirez-Vazquez, R., Escobar, I., Vandenbosch, G. A. E., & Arribas, E. (2024). Personal exposure to radiofrequency electromagnetic fields: A comparative analysis of international, national, and regional guidelines. Environmental Research, 246, 118124. https://doi.org/10.1016/j.envres.2024.118124

- Rangan, S., Rappaport, T. S., & Erkip, E. (2014). Millimeter-Wave Cellular Wireless Networks: Potentials and Challenges. Proceedings of the IEEE, 102(3), 366–385. https://doi.org/10.1109/JPROC.2014.2299397

- Sagar, S., Dongus, S., Schoeni, A., Roser, K., Eeftens, M., Struchen, B., Foerster, M., Meier, N., Adem, S., & Röösli, M. (2018). Radiofrequency electromagnetic field exposure in everyday microenvironments in Europe: A systematic literature review. Journal of Exposure Science & Environmental Epidemiology, 28(2), 147–160. https://doi.org/10.1038/jes.2017.13

- Sandoval-Diez. (2024). Determining the relationship between mobile ... | Open Research Europe. https://open-research-europe.ec.europa.eu/articles/4-206

- Sim, C.-Y.-D., Liu, H.-Y., & Huang, C.-J. (2020). Wideband MIMO Antenna Array Design for Future Mobile Devices Operating in the 5G NR Frequency Bands n77/n78/n79 and LTE Band 46. IEEE Antennas and Wireless Propagation Letters, 19(1), 74–78. https://doi.org/10.1109/LAWP.2019.2953334

- Statistiek, C. B. voor de. (n.d.). Wijk- en buurtkaart 2024 [Webpagina]. Centraal Bureau voor de Statistiek. Retrieved 31 March 2025, from https://www.cbs.nl/nl-nl/dossier/nederland-regionaal/geografische-data/wijk-en-buurtkaart-2024

- Traini, E., Martens, A. L., Slottje, P., Vermeulen, R. C. H., & Huss, A. (2023). Time course of health complaints attributed to RF-EMF exposure and predictors of electromagnetic hypersensitivity over 10 years in a prospective cohort of Dutch adults. Science of The Total Environment, 856, 159240. https://doi.org/10.1016/j.scitotenv.2022.159240

- Urbinello, D., Joseph, W., Huss, A., Verloock, L., Beekhuizen, J., Vermeulen, R., Martens, L., & Röösli, M. (2014). Radio-frequency electromagnetic field (RF-EMF) exposure levels in different European outdoor urban environments in comparison with regulatory limits. Environment International, 68, 49–54. https://doi.org/10.1016/j.envint.2014.03.007

- Urbinello, D., Joseph, W., Verloock, L., Martens, L., & Röösli, M. (2014). Temporal trends of radio-frequency electromagnetic field (RF-EMF) exposure in everyday environments across European cities. Environmental Research, 134, 134–142. https://doi.org/10.1016/j.envres.2014.07.003

- Velghe, M., Aerts, S., Martens, L., Joseph, W., & Thielens, A. (2021). Protocol for personal RF-EMF exposure measurement studies in 5th generation telecommunication networks. Environmental Health, 20(1), 36. https://doi.org/10.1186/s12940-021-00719-w

- Velghe, M., Joseph, W., Debouvere, S., Aminzadeh, R., Martens, L., & Thielens, A. (2019). Characterisation of spatial and temporal variability of RF-EMF exposure levels in urban environments in Flanders, Belgium. Environmental Research, 175, 351–366. https://doi.org/10.1016/j.envres.2019.05.027

- Wali, S. Q., Sali, A., Allami, J. K., & Osman, A. F. (2022). RF-EMF Exposure Measurement for 5G Over Mm-Wave Base Station With MIMO Antenna. IEEE Access, 10, 9048–9058. https://doi.org/10.1109/ACCESS.2022.3143805

# Appendix

Measurement Details and Procedure

A.1 Measurement details

A1.1 Static Measurements

1. Amsterdam City Centre

- Start time: 09:30 16/12/24

- End Time: 10:00 17/12/24

- Setting: Urban

- Environment: Outdoor

2. Utrecht City Centre

- Start time: 09:56 24/11/24

- End Time: 13:00 25/11/24

- Setting: Urban

- Environment: Outdoor

3. Rotterdam City Centre

- Start time: 10:25 06/11/24

- End Time: 11:14 07/11/24

- Setting: Urban

- Environment: Outdoor

4. Rotterdam City Centre

- Start time: 10:45 13/11/24

- End Time: 10:45 14/11/24

- Setting: Urban

- Environment: Indoor

5. Doorn

- Start time: 10:40 05/12/24

- End Time: 11:00 06/12/24

- Setting: Rural

6. Puttershoek

- Start time: 11:15 10/11/24

- End Time: 11:15 11/11/24

- Setting: Rural

- Environment: Outdoor

7. Baambrugge

- Start time: 10:56 19/11/24

- End Time: 12:00 20/11/24

- Setting: Rural

- Environment: Outdoor

8. Baambrugge

- Start time: 09:58 26/11/24

- End Time: 11:00 27/11/24

- Setting: Rural

- Environment: Indoor

A.1.2 Mobile Measurements

1. Utrecht University Library City Centre

- Start time: 09:38 10/12/24

- End Time: 09:54 10/12/24

- Setting: Urban

2. Utrecht University Library Uithof

- Start time: 10:51 12/12/24

- End Time: 11:07 12/12/24

- Setting: Urban

3. Plus Romijn Kockengen

- Start time: 10:53 11/12/24

- End Time: 11:08 11/12/24

- Setting: Rural

4. Spar Cothen

- Start time: 10:01 11/12/24

- End Time: 10:16 11/12/24

- Setting: Rural

A.2 Measurements steps

A.2.1 Preparation

- Ensure that all three phones are fully charged.

- Equip the phones with the prepaid SIM cards, and verify they have sufficient credit.

- When starting measurements ensure all the phones are on screen lock, connected to the mobile data network (WIFI and Bluetooth off).

- Open the ETAIN 5G Scientist app and press start.

- Set up the ExpoM-RF 4 and make sure it is fully charged.

- Turn on the ExpoM-RF 4 and make sure a red-light is flashing, indicating it is measuring.

A.2.2.1 Static Measurements

- Start the measurement between 9:30 and 11:30 am CET.

- Change the measurement interval of the ExpoM-RF 4 to 12 seconds and disable GPS coordinates

- Setup the phones and ExpoM-RF 4 as illustrated in Figure 2.3.1.

- For each measurement, ensure that no large conducting objects or surfaces are in proximity (within 0,5 m) of the experimental setup. Which could result in reflection of the RF-EMF, causing inaccurate measurements by the ExpoM-RF 4 (Korkmaz et al., 2024).

- Allow the measurement to run for at least 24 hours continuously.

A.2.2.2 Mobile Measurements

- Set the measurement interval of the ExpoM-RF 4 to 7 seconds and enable GPS coordinates.

- Place the three phones and the ExpoM-RF 4 as high and as far from the human body as possible in the backpack. This to prevent body shielding causing inaccurate readings by the ExpoM-RF 4.

- Walk very slowly 7.5 minutes Outdoor around the Indoor location and move then Indoor. Walk here for another 7.5 minutes, stay close to the outdoor wall.

A.2.3 Data Collection

- Collect the data from the ETAIN 5G Scientist app’s data base and save as CSV file.

- Use the Fields at Work utility software to collect the ExpoM-RF 4 data and save as CSV file.

- Store all data in organized folders with a consistent naming format, so data is easily findable for analysis in RStudio.

A.2.4. Documentation during each measurement

- Start and end time/data

- Location, GPS coordinates and address

- Description of the measurement setting

- Protocol deviations

Conversion Data ETAIN 5G Scientist app

B.1 Data Conversion steps

B.1.1 App Calibration function

Use the conversion and far field functions to convert the dBm values from the ETAIN 5G Scientist app towards the estimated far field downlink exposure in power flux density (mWm2).

- Extract frequencies from channel numbers.

- Normalize the RSRP values to 1800 MHz for each cell, Registered and Neighboring.

- Apply app’s calibration function.

- Convert decibels micro volts per meter (dBμV/m) to V/m.

- Convert V/m to Power Flux density (mWm2).

- Combine registered and all neighboring cells together to get the estimated exposure for each provider.

B.1.2 Mapping Calibration function

Use the conversion and far field functions to convert the dBm values from the ETAIN 5G Scientist app towards the estimated environmental downlink exposure in power flux density (mWm2).

- Convert all RSRP values (Registered and Neighboring cells) from dBm to mW.

- Sum up all cells RSRP values for each measurement.

- Create 10s (Mobile) and 1 min (Static) time intervals from start time.

- Calculate the median for each provider (KPN median + Odido median + Vodafone median).

- Convert all values back to dBm.

- Use Mapping Calibration function to calculate the values in dBμV/m.

- Convert from dBμV/m to V/m to Power flux density (mWm2).

B.1.3 ExpoM-RF 4

ExpoM-RF 4 data is measured in V/m, it is only needed to convert this to power flux density (mWm2) and add the different frequencies to get the total downlink exposure.

- Convert the ExpoM-RF 4 data which is measured in V/m also towards power flux density (mWm2) and calculate the total downlink Exposure for LTE.

- Total downlink exposure is calculated as the sum of the mobile frequencies used by LTE: 700, 800, 900, 1800, 2100 and 2600 MHz (Sandoval-Diez, 2024).

B.2 Conversion Formulas

1. Normalize RSRP values to the reference frequency (f0) :1800 Mhz

normalizedSSIs= RSRP_measured + 20 * log10(frequencies / f0)

2. App’s calibration function:

Exposure (dBμV/m) = 154.74453 + 0.61062*sLTR_RSRP

3. Convert decibels micro volts per meter (dBμV/m) to V/m

Exposure (V/m) = 10^(Exposure (dBµV/m) / 20) / 106

4. Convert V/m to Power Flux density (mWm2)

Exposure (mWm2) = Exposure (V/m) * 103/ 3767

5. Convert all RSRP values from dBm to mW

Exposure (mW) = 10^(Exposure (dBm) / 10)

6. Convert RSRP values from mW to dBm

Exposure (dBm) = 10 * log10(Exposure (mW)

7. Mapping Calibration function

Exposure (dBμV/m) = 168.442 + 0.75985 * aLTE_RSRP

Variability analysis for each location

C.1 Variability for each static location

C.1.1 Amsterdam City Centre

Comparison between Timeframes

Table C1.1.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.1.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.1.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.1.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.1.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C.1.2 Utrecht City Centre

Comparison between Timeframes

Table C1.2.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.2.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.2.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.2.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.2.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C.1.3 Rotterdam City Centre (Outdoor)

Comparison between Timeframes

Table C1.3.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.3.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.3.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.3.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.3.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C.1.4 Rotterdam City Centre (Indoor)

Comparison between Timeframes

Table C1.4.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.4.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.4.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.4.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.4.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C1.5 Doorn

Table C1.5.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.5.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.5.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.5.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.5.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C1.6 Puttershoek

Comparison between Timeframes

Table C1.6.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.6.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.6.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.6.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.6.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C1.7 Baambrugge (Outdoor)

Comparison between Timeframes

Table C1.7.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.7.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.7.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.7.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.7.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C1.8 Baambrugge (Indoor)

Comparison between Timeframes

Table C1.8.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across different timeframes and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C1.8.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone).

Table C1.8.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure C1.8.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes.

Provider Comparison

Figure C1.8.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each timeframe is indicated with: **** = p < 0.0001.

C.2 Variability for each mobile location

C.2.1 Utrecht University Library City Centre

Indoor - Outdoor

Table C2.1.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C2.1.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between environments is indicated with: **** = p < 0.0001.

Table C2.1.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, MAD, and IQR.

Figure C2.1.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between conditions is indicated with: **** = p < 0.0001.

Provider Comparison

Figure C2.1.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each location is indicated with: **** = p < 0.0001 and ‘ns’ = not significant.

C.2.2 Utrecht University Library Uithof

Indoor - Outdoor

Table C2.2.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C2.2.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between environments is indicated with: **** = p < 0.0001.

Table C2.2.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, MAD, and IQR.

Figure C2.2.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between conditions is indicated with: **** = p < 0.0001.

Provider Comparison

Figure C2.2.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each location is indicated with: **** = p < 0.0001, *** = p < 0.001.

C.2.3 Plus Romijn Kockengen

Indoor - Outdoor

Table C2.3.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C2.3.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between environments is indicated with: **** = p < 0.0001.

Table C2.3.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, MAD, and IQR.

Figure C2.3.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between conditions is indicated with: **** = p < 0.0001.

Provider Comparison

Figure C2.3.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each location is indicated with: * = p < 0.05, and ‘ns’ = not significant.

C2.4 Spar Cothen

Indoor - Outdoor

Table C2.4.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Figure C2.4.1 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between environments is indicated with: **** = p < 0.0001.

Table C2.4.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, MAD, and IQR.

Figure C2.4.2 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between conditions is indicated with: **** = p < 0.0001.

Provider Comparison

Figure C2.4.3 Comparison of downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations and three mobile network providers (KPN, Odido, Vodafone). Statistical significance between providers within each location is indicated with: **** = p < 0.0001, *** = p < 0.001, and ‘ns’ = not significant.

Uncertainty analysis for each location

D.1. Uncertainty for each static location

D.1.1 Amsterdam City Centre

Table D1.1.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.1.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.1.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak negative correlation was observed (R = –0.093, p = 0.00035)

D.1.2 Utrecht City Centre

Table D1.2.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.2.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.2.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak positive correlation was observed (R = 0.075 , p = 0.0036)

D.1.3 Rotterdam City Centre (Outdoor)

Table D1.3.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.3.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.3.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant moderate positive correlation was observed (R = 0.48 , p < 2.2e−16).

D.1.4 Rotterdam City Centre (Indoor)

Table D1.4.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.4.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.4.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak positive correlation was observed (R = 0.16 , p = 2.8e−10).

D.1.5 Doorn

Table D1.5.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.5.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.5.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak positive correlation was observed (R = 0.16 , p = 7.8e−16).

D.1.6 Puttershoek

Table D1.6.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.6.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.6.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak negative correlation was observed (R = -0.18 , p = 1.3e−11).

D.1.7 Baambrugge (Outdoor)

Table D1.7.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.7.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.7.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak negative correlation was observed (R = -0.22 , p < 2.2e−16).

D.1.8 Baambrugge (Indoor)

Table D1.8.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Reported values include the median, MAD, and IQR.

Figure D1.8.1 Comparison of total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes. Statistical significance between measurement sources within each timeframe is indicated with: **** = p < 0.0001.

Figure D1.8.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak positive correlation was observed (R = 0.29 , p < 2.2e−16).

D.2. Uncertainty for each mobile location

D.2.1 Utrecht University Library City Centre

Figure D2.1.1 Comparison of total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between measurement sources is indicated with: *** = p < 0.001 and ‘ns’ = not significant.

Table D2.1.1 Descriptive statistics for total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, IQR, and MAD.

Figure D2.1.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant weak positive correlation was observed (R = 0.21 , p = 0.038).

D.2.2 Utrecht University Library Uithof

Figure D2.2.1 Comparison of total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between measurement sources is indicated with: **** = p < 0.0001 and ‘ns’ = not significant.

Table D2.2.1 Descriptive statistics for total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, IQR, and MAD.

Figure D2.2.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant strong positive correlation was observed (R = 0.83 , p < 2.2e−16).

D.2.3 Plus Romijn Kockengen

Figure D2.3.1 Comparison of total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations. Statistical significance between measurement sources is indicated with: *** = p < 0.001 and ‘ns’ = not significant.

Table D2.3.1 Descriptive statistics for total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations. Reported values include the median, IQR, and MAD.

Figure D2.3.2 Correlation between total downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4. A significant moderate positive correlation was observed (R = 0.66 , p < 2.2e−16).

Tables comparison Static and Mobile measurements

E.1. Variability analysis Static

Table E1.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations, settings (rural vs. urban), and mobile network providers (KPN, Odido, Vodafone). Reported values include the median, IQR, and MAD.

Table E1.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations and settings (rural vs. urban). Reported values include the median, MAD, and IQR.

Table E1.3 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across six timeframes, two settings (rural vs. urban), and three mobile network providers (KPN, Odido, Vodafone). Reported values include the median, MAD, and IQR.

Table E1.4 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across six timeframes and two settings (rural vs. urban). Reported values include the median, MAD, and IQR.

E.2. Variability analysis Mobile

Table E2.1 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App across indoor and outdoor locations, settings (rural vs. urban), and mobile network providers (KPN, Odido, Vodafone). Reported values include the median, IQR, and MAD.

Table E2.2 Descriptive statistics for downlink RF-EMF exposure (mW/m²) measured by the ExpoM-RF 4 across indoor and outdoor locations and settings (rural vs. urban). Reported values include the median, IQR, and MAD.

E.3. Uncertainty Analysis Static

Table E3.1 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations and settings (rural vs. urban). Reported values include the median, IQR, and MAD.

Table E3.2 Descriptive statistics for total environmental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across six timeframes and two settings (rural vs. urban). Reported values include the median, IQR, and MAD.

E.4. Uncertainty Analysis Mobile

Table E4.1 Descriptive statistics for total environemental downlink RF-EMF exposure (mW/m²) measured by the ETAIN 5G Scientist App and the ExpoM-RF 4 across indoor and outdoor locations and settings (rural vs. urban). Reported values include the median, IQR, and MAD.